PFRC-2 X-ray report from the run of 2020/08/14

Prepared 2020/08/18

Charles Swanson

Machine configuration:

Inner 4x L-2-B set coils on each side powered by 2x TSD60-832/480+HS Magna-Power power supplies

Outer 4x L-2-B set coils on each side powered by “Big Blue,” Kinetics Control Systems Model SVRR-0125PM42TSY

L-2-A set coils on each side powered by “Big Blue,” Kinetics Control Systems Model SVRR-0125PM42TSY

Nozzle coils powered by rollaround power supply

SEC pumped by Leybold-Heraeus TurboVac 450, closed

CC pumped by 100 l/s Varian TV 141 Navigator, open

FEC pumped by Leybold-Heraeus TurboVac 450, open

Seed plasma RF frequency 27 MHz

RMF frequency 4.3 MHz

Room temperature flux conservers

H2 gas feed into SEC

X-ray configuration:

SDD3, serial number 19777, in CC midpoint mount installed on 2016/08/30 and whose variable apertures were changed on 2019/02/21

SDD1, serial number 27054 after repairs, in CC radial scan mount installed on 2018/01/24

Run parameter envelope:

Magna-Power current 170 - 280 A

Big Blue current 166 - 345 A

Nozzle current 77 - 101 A

CC gas pressure 0.466 – 0.779 mTorr

Seed forward power 0 - 43 W, CW

RMF forward power 17 – 53 kW

Experiments:

RTD on seed plasma

X-ray time trace during density fall and steady state

RTD on RMF plasma

X-ray bursts during densification

Systematic X-ray signal during densification

Summary:

The SDD’s Rise Time Discrimination (RTD) feature requires a clear signal in the fast channel, and so will be of only limited use when the fast channel is also experiencing pileup.

A condition was observed during which the piled-up count rate (probably monotonic with UV flux) had a peak when the density was falling but not yet at its minimum.

There are two X-ray behaviors observed during densification: At lower RMF power, field, and 0 seed power, a very stochastic “bursting” phenomenon exists in which hundreds of RMF pulses yield no counts, then a single RMF pulse will produce several.

At higher RMF power and field, a systematic X-ray signal approximately 250 μs after the RMF pulse turns on is measured. This is approximately when densification occurs.

Recommended future work:

The systematic X-rays at 250 μs: Are they single-counts? Are they piled up? An aperture scan at these conditions would tell us.

Experiment: RTD on seed plasma:

Used the Rise Time Discrimination (RTD) feature on the SDD during measurement of seed plasma X-ray spectrum

Started 9:40 AM, No Spectrum was recorded

Ended 9:44 AM

Experiment parameter envelope:

Magna-Power current 170 A

Big Blue current 170 A

Nozzle current 77 A

CC gas pressure 0.466 mTorr

Seed forward power 43 W, CW

RMF forward power 0 kW

Figures:

No figures

Numerical results:

RTD off: 7 counts/s total

RTD on, fast threshold set very high: No counts!

RTD on, fast threshold correctly set: Count rate agrees with RTD off

Comments:

It is apparent that the start time recorded by the SDD is from the fast channel, while the peak time is from the slow channel. RTD uses both the fast and slow channels, and determines the rise time from the discrepancy in timing. This also explains a phenomenon seen on 2020/07/31 in the X-ray calibration tube: The deficit at low energy when RTD was on. Counts with less energy than the fast threshold were suppressed.

RTD is therefore not expected to help us mitigate pulse pileup operationally in the PFRC-2.

Experiment: X-ray time trace during density fall and steady state:

At these conditions, the interferometer (density) was observed to first fall, then become steady. The X-ray (pilup, UV) signal was observed to peak during the end of the fall, then again during the steady-state.

Started 9:57 AM, Spectra 07,08

Ended 10:00 AM

Experiment parameter envelope:

Magna-Power current 174 A

Big Blue current 166 A

Nozzle current 80 A

CC gas pressure 0.497 mTorr

Seed forward power 43 W, CW

RMF forward power 41.3 kW

Figures:

At left: Red: X-channel of interferometer, single shot. Blue: Y-channel of interferometer, single shot. Green: Average of arctangent of Y and X.

At right: Yellow: RMF pulse. Red: Single shot (different from left) X-rays from SDD3 in CC midpoint mount. Pink: Persistent histogram of these X-rays. Blue: Single shot X-rays from SDD1 in CC radial scan mount. Light blue: Persistent histogram of these X-rays.

Above: This is Spectrum 07, SDD3 in midpoint mount with Aperture 3 in place. It follows the pattern of an “anomalous” or “no-lines-mode” spectrum, attributed to pulse pileup. The count rate is supposed to be monotonic with UV flux, but not necessarily linear.

Numerical results:

Interferometer (density) spends the first 3.5 ms falling from 1.1 to 0.5 radians, then spends the last 3.5 ms steadier falling only from 0.5 radians to 0.4 radians. Assuming a 5 cm radius, this would correspond to ne=9×1011, 4.1×1011, 3.3×1011/cc respectively.

Both X-ray detectors peak at roughly 2.7 ms, have a minimum at roughly 3.7 ms, then rise to a flat top at roughly 6 ms.

Comments:

The first peak occurs when the density is falling. It appears that the photon (X-ray, UV) signal is correlated with a decrease in density. Perhaps the mechanism for this density loss is particle loss to the walls, which then causes photon emission.

The second lobe occurs when the density is at a steady, low value. This has been observed other times. The photon (X-ray, UV) signal is highest when the density signal is lowest.

Experiment: RTD on RMF plasma:

Various combinations of fast peaking time, slow peaking time, fast threshold, show threshold, and RTD were tried.

Started 10:00 AM, no Spectra recorded

Ended 10:55 AM

Experiment parameter envelope:

Magna-Power current 174 A

Big Blue current 166 A

Nozzle current 80 A

CC gas pressure 0.497 mTorr

Seed forward power 43 W, CW

RMF forward power 41.3 kW

Figures:

None

Numerical results:

| Slow threshold (channel) | Fast threshold (channel) | Fast risetime (ns) | Slow risetime (ns) | RTD sensitivity (%) | Fast channel rate (c/s) | Show channel rate (c/s) |
| --- | --- | --- | --- | --- | --- | --- |
| 32 | 32 | 200 | 200 | No | 60 | 3500 |
| 32 | 25 | 200 | 200 | No | 900 | 3500 |
| 32 | 22 | 200 | 200 | No | 3200 | 3800 |
| 32 | 22 | 200 | 200 | 120 | 3000 | 2700 |
| 0 | 22 | 200 | 200 | 120 | 3100 | 5000 |
| 0 | 22 | 200 | 400 | 120 | 3000 | 2800 |
| 32 | 22 | 200 | 400 | No | 2400 | 24000 |
| 32 | 22 | 200 | 1000 | No | 2500 | 25000 |
| 32 | 22 | 200 | 1000 | 120 | 2400 | 100 |
| 32 | 22 | 200 | 1000 | 120 | 2400 | <10 |
| 32 | 22 | 200 | 1000 | No | 2500 | 25000 |

Comments:

When fast and slow channel peaking times disagree, RTD can have a sharply inhibitory effect. This is presumably due to pileup resulting in apparent photons at different times in the different channels. This might be a good thing, as pileup is inhibited. But it is also likely that pileup will obscure true X-rays, so the true signal will also be inhibited.

Experiment: X-ray bursts during densification:

On previous days (2020/03/05, 06, and 09) we had seen very intermittent “bursts” of counts at the beginning (<0.5ms) of a discharge. Hundreds of thousands of RMF pulses would go by with no counts, then one RMF pulse would have several. We were not able to achieve this condition today until we turned off the seed plasma.

Started 10:49 AM, Spectra 10-16

Ended 12:00 PM

Experiment parameter envelope:

Magna-Power current 174 - 200 A

Big Blue current 166 - 243 A

Nozzle current 80 - 101 A

CC gas pressure 0.497 - 0.779 mTorr

Seed forward power 0 - 43 W, CW (achieved bursts only at 0 W)

RMF forward power 33 - 44.6 kW

Figures:

Spectrum and time-trace of a burst, not of the same burst. Left: Burst counts are of low energy, <600 eV. Right: Burst counts come in early in the discharge, <0.5 ms.

Numerical results:

In 8 minutes of operation, only 4 RMF pulses produced burst counts. RMF pulses occurred 2x per second. That’s 1 pulse of every 240 producing bursts.

Comments:

The cause is unknown. Perhaps the densification mechanism relies on a wall interaction, and the column hits the wall at a different azimuthal position each time. Perhaps these are not true photon counts.

Experiment: Systematic X-ray signal during densification:

By going to higher power and magnetic field, we were able to observe X-rays coming in systematically at some time during the densification, usually around 0.25 ms. This corresponds well to the densification.

Started 12:10 PM, Spectra 17 - 25

Ended 12:30 PM

Experiment parameter envelope:

Magna-Power current 280 A

Big Blue current 345 A

Nozzle current 101 A

CC gas pressure 0.677 mTorr

Seed forward power 0 - 110 W, CW

RMF forward power 53 kW

Figures:

Seed power 0W:

Seed power 50 W:

Seed power 110 W:

Seed power 10 W:

Numerical results:

| Seed power | X-ray count rate | X-ray average energy E-293 eV | Time of max emission |
| --- | --- | --- | --- |
| 0 W | 2970±180 c/s | 103±8 eV | 270 μs |
| 50 W | 2360±110 c/s | 133±7 eV | 250 μs |
| 110 W | 2430±150 c/s | 144±10 eV | 280 μs |
| 10 W | 2950±210 c/s | 154±13 eV | 230 μs |

Comments:

Density during these conditions had a sudden rise between 200 μs and 270 μs. The tail end of this period is when the X-rays came in. The seed power did not change measurements appreciably. It does appear that there is a systematic tendency for X-rays to be measured during densification, supporting the hypothesis that a population of fast electrons is essential to this process.

Experiment: Name:

Summary

Started, Spectra

Ended

Experiment parameter envelope:

Figures:

Numerical results:

Comments: