PFRC-2 X-ray report from the run of 2020/09/25

Prepared 2020/10/19

Charles Swanson

Machine configuration:

Inner 4x L-2-B set coils on each side powered by 2x TSD60-832/480+HS Magna-Power power supplies

Outer 4x L-2-B set coils on each side powered by “Big Blue,” Kinetics Control Systems Model SVRR-0125PM42TSY

L-2-A set coils on each side powered by “Big Blue,” Kinetics Control Systems Model SVRR-0125PM42TSY

Nozzle coils powered by rollaround power supply

SEC pumped by Leybold-Heraeus TurboVac 450, closed

CC pumped by 100 l/s Varian TV 141 Navigator, open

FEC pumped by Leybold-Heraeus TurboVac 450, open

Seed plasma RF frequency 27 MHz

RMF frequency 4.3 MHz

Room temperature flux conservers

Kr gas feed into SEC

X-ray configuration:

SDD3, serial number 19777, in CC midpoint mount installed on 2016/08/30, whose variable apertures were changed on 2019/02/21, with the 850 nm removable Mylar filter installed on 2020/09/17

SDD1, serial number 27054 after repairs, in CC radial scan mount installed on 2018/01/24

SDD2, serial number 16645, in SEC long, close aperture mount installed on 2020/09/24

Run parameter envelope:

Magna-Power current 141 - 250 A

Big Blue current 170 - 369 A

Nozzle current 55 A

CC gas pressure 0.357 – 0.485 mTorr

Seed forward power 35 - 240 W, CW

RMF forward power 0, 60 kW

Experiments:

Radial scan of SEC seed plasma

SEC seed comparison, bright blue vs dim purple spectral comparison

Mylar filter seed plasma comparison

Strange SDD effect: Lines at shifted energies

Strange SDD effect: Nontrivial Slow Threshold behavior

Diamagnetic Loop Shield grounded/floating comparison

Summary:

During the seed plasma, the diameter of the column in the SEC is 3.5 cm. This is the impact parameter at which the integrated count rate drops to half of maximum.

During the seed plasma, the plasma can appear bright and blue (presumed helicon) or dim and purple (presumed capacitive). The count rate and average energy of the dim purple mode are higher.

The Mylar filter window behaves as expected during seed plasma operation

Sometimes, the C K-α and O K-α lines appear with some offset energy. I speculate that this is due to RMF noise moving the baseline of the detector, as shielding mitigates this phenomenon.

As with yesterday’s run, There’s an odd slow threshold behavior in which setting a small but finite slow threshold reveals a distinct peak that was not visible at slow threshold zero. I speculate that this is due to RMF noise moving the baseline of the detector.

When DL1 shield is grounded, the X-ray count rate decreases rather than increasing! This is counter to what would be expected from wall emission.

Recommended future work:

Add Mylar filter to CC radial scan mount, take radial scan of X-ray spectrum at a couple B field values

Experiment: Radial scan of SEC seed plasma:

During the dim purple mode, the aperture slit on the SEC long detector mount (with close aperture) was swept to produce a radial profile of the count rate. The count rate drifted over time, so this was corrected-for in the count rate plot. The column decays to half count rate at a radius of 1.7 cm. There was a finite count rate even when the detector aperture was pointing at the ID of the stainless steel pipe.

Started 10:52 AM, Spectra 09 - 17

Ended 11:10 AM

Experiment parameter envelope:

Magna-Power current 200 A

Big Blue current 178 A

Nozzle current 55 A

CC gas pressure 0.485 mTorr

Seed forward power 240 W, CW

RMF forward power 0 kW

Figures:

Left: Count rate at r=0 was measured twice, and found to vary by 40%. It was assumed that a multiplicative factor varied linearly with time, and this was corrected-for in the plot. The red line marks the impact parameter outside of which the SDD aperture was looking at the inner diameter of the mount tube.

Right: Spectrum from 4 cm impact parameter, looking at the ID of the tube mount and not the plasma. There are still 0.66 c/s count rate, and a broad spectrum!

Comments:

There must be X-rays from the plasma scattering into the tube. The plasma column is around 3.5 cm in diameter.

Experiment: SEC seed comparison, bright blue vs dim purple spectral comparison:

The SEC seed plasma could appear to be a bright blue column (presumed to be helicon) or a dim purple column (presumed to be capacitive). The bright blue column had lower count rate and lower average X-ray energy (shaper, cooler).

Started 11:19 AM, Spectra 21 and 23

Ended 11:21 AM

Experiment parameter envelope:

Magna-Power current 200 A

Big Blue current 178 A

Nozzle current 55 A

CC gas pressure 0.485 mTorr

Seed forward power 240 W, CW

RMF forward power 0 kW

Figures:

Blue: Dim purple mode. Red: Bright blue mode.

Numerical results:

| Spectrum | Condition | Count rate | Average X-ray energy above 200 eV |
| --- | --- | --- | --- |
| 21 | Dim purple | 910±10 c/s | 289±3 eV |
| 23 | Bright blue | 18±2 c/s | 109±9 eV |

Comments:

The mode presumed to be capacitive is certainly much hotter and denser than the mode presumed to be helicon.

Experiment: Mylar filter seed plasma comparison:

The Mylar filter perturbs the seed plasma in pretty much the expected fashion

Started 11:40 AM, Spectra 32 and 34

Ended 11:54 AM

Experiment parameter envelope:

Magna-Power current 250 A

Big Blue current 369 A

Nozzle current 55 A

CC gas pressure 0.357 mTorr

Seed forward power 215 W, CW

RMF forward power 0 kW

Figures:

Left: Raw spectra. Blue: No filter. Red: Mylar filter in place.

Right: Same, except red has been divided by the transmission at that energy.

Comments:

Except for the behavior at spectral lines, the blue and corrected red curves lie mostly atop each other. As has been discussed before in these reports, the transmission is not expected to be correct at spectral lines, as the true width of spectral lines is much smaller than the detector resolution.

Experiment: Strange SDD effect: Lines at shifted energies:

As with yesterday, sometimes the C K-α and O K-α lines appeared roughly correct, but with an offset to the energy. This run was in better agreement with the predicted values than yesterday’s run.

Started 12:07 PM, Spectra 36 - 67

Ended 1:47 PM

Experiment parameter envelope:

Magna-Power current 141 A

Big Blue current 185 - 187 A

Nozzle current 55 A

CC gas pressure 0.381 – 0.441 mTorr

Seed forward power 35 W, CW

RMF forward power 60 kW

Figures:

Example fit, showing the shifted spectral lines

Numerical results:

Spectrum 36 as an example:

| Line | Expected | Measure fit | Discrepancy |
| --- | --- | --- | --- |
| C K-α | 277 eV | 294 eV | 17 eV |
| O K-α | 525 eV | 537 eV | 12 eV |

Comments:

I speculate that this effect is due to RMF noise coupling into the SDD and moving the baseline, as shielding the SDD improved the discrepancy.

Experiment: Strange SDD effect: Nontrivial Slow Threshold behavior:

Typically, setting the SDD’s “Slow Threshold” just serves to reject X-rays with measured energies below a certain energy. It shouldn’t distort the spectrum other than to impose a sharp cutoff. However, in today’s and yesterday’s runs, it was observed that setting Slow Threshold < channel 5 distorted the spectra!

Started 1:04 PM, Spectra 48 - 54

Ended 1:15 PM

Experiment parameter envelope:

Magna-Power current 141 A

Big Blue current 187 A

Nozzle current 55 A

CC gas pressure 0.441 mTorr

Seed forward power 35 W, CW

RMF forward power 60 kW

Figures:

Spectra with several Slow Thresholds.

Yellow: Ch23, typical.

Red: Ch5, low enough to let the thermal and electrical detector noise in. This is typical behavior; the spectrum above ch23 is identical to the yellow spectrum.

Blue: Ch0, the lowest possible value. The entire spectrum is distorted! A peak visible in the other two is covered by anomalous counts.

Comments:

I don’t know why this is happening, but I speculate that it is RMF noise coupling into the detector. During seed plasma operation, this has historically worked as expected. In the future, we must verify whether there are well-defined peaks using ST5 or higher, rather than ST0.

Experiment: Diamagnetic Loop Shield grounded/floating comparison:

The shield of DL1 was floating for the majority of these runs. When it was grounded, the count rate decreased, contrary to what would be expected from wall emission.

Started 1:45 PM, Spectra 66 - 67

Ended 1:49 PM

Experiment parameter envelope:

Magna-Power current 141 A

Big Blue current 187 A

Nozzle current 55 A

CC gas pressure 0.441 mTorr

Seed forward power 35 W, CW

RMF forward power 60 kW

Figures:

Numerical results:

| Spectrum | Condition | Count rate | Average X-ray energy above 200 eV |
| --- | --- | --- | --- |
| 66 | Floating | 3540±60 c/s | 128±2 eV |
| 67 | Grounded | 1700±40 c/s | 124±3 eV |

Comments:

I expected that this midpoint mount spectrum was due to wall emission rather than volumetric emission. The midpoint mount looks at some Aluminum foil which is electrically shorted to the DL1 shield, so grounding it should increase the electron current to it, increasing the X-ray count rate. However this is not what happened.

Experiment: Name:

Summary

Started, Spectra

Ended

Experiment parameter envelope:

Figures:

Numerical results:

Comments: